MRI-Guided Microwave Thermal Ablation

Formulation as a PDE-Constrained Optimal Control Problem

Technical Report — Free Final Time, State-Constrained Variational Formulation

1. Problem Overview and Physical Setting

MRI-guided thermal ablation is the clinical procedure of using a controlled energy source (microwave, radiofrequency, laser, or focused ultrasound) to heat and destroy a tumor target while sparing surrounding healthy tissue. Real-time MRI temperature imaging (MR thermometry, via the proton resonance frequency shift) provides online feedback of the three-dimensional temperature field, making this a natural setting for closed-loop, feedback-optimal control.

We denote the tissue domain by Ω ⊂ ℝ³ (or the cylindrical domain Ω ⊂ ℝ² in the axisymmetric (r, z) case). The boundary is ∂Ω, with inward unit normal n̂. The time horizon is t ∈ [0, t_f], where the final time t_f > 0 is a free optimization variable.

1.1 Spatial Coordinates

Cartesian: r = (x, y, z) ∈ Ω ⊂ ℝ³

Cylindrical (axisymmetric): r = (r, θ, z)

MRI voxel index (discretized): i ∈ {1, 2, …, N},  N = N_x × N_y × N_z

2. State Variables

The system state is a pair of spatially-distributed fields:

2.1 Temperature Field

T = T(r, t) : Ω × [0, t_f] → ℝ  [°C or K]

T(r, t) is the primary observable — it is measured in real time by MRI thermometry. In the discretized system, the temperature state vector is:

x_T(t) = [T₁(t), T₂(t), …, T_N(t)]ᵀ ∈ ℝᴺ

where Tᵢ(t) = T(rᵢ, t) is the temperature at the i-th voxel centroid rᵢ

2.2 Thermal Damage Field (Arrhenius)

Ω_d = Ω_d(r, t) : Ω × [0, t_f] → ℝ₊  [dimensionless]

Ω_d accumulates irreversible tissue damage. It is defined by the Arrhenius thermal dose model:

Ω_d(r, t) = ∫₀ᵗ A · exp(−E_a / (R · T(r, τ))) dτ

Equivalently in ODE form:

dΩ_d/dt = A · exp(−E_a / (R · T)),    Ω_d(r, 0) = 0

where A [s⁻¹] is the frequency factor, E_a [J/mol] is the activation energy, and R = 8.314 J/(mol·K) is the universal gas constant. The threshold Ω_d = 1 corresponds to 63% probability of cell death (irreversible necrosis).

2.3 Combined Distributed State

The complete infinite-dimensional state is:

𝒳(r, t) = (T(r, t), Ω_d(r, t)),    r ∈ Ω,  t ∈ [0, t_f]

After spatial discretization, the lumped finite-dimensional state vector is:

x(t) = [x_T(t)ᵀ, x_Ω(t)ᵀ]ᵀ ∈ ℝ²ᴺ

3. Control Variables

The control input depends on the ablation modality. We consider the general microwave/RF case:

3.1 Scalar Power Control (Single Applicator)

u(t) = P(t) ∈ [0, P_max]  [W]

3.2 Vector Control (Robotic MRI-Guided Ablation)

For a robotic system with actuated probe position and orientation:

u(t) = [P(t), q(t)]ᵀ ∈ ℝ^(1+n_q)

where q(t) ∈ ℝⁿq is the robot joint configuration vector.

3.3 Control Constraints

u(t) ∈ 𝒰 := {u : 0 ≤ P(t) ≤ P_max, |u̇(t)| ≤ u̇_max}

The control set 𝒰 encodes physical actuator limits (maximum power, maximum slew rate, probe mechanical limits).

4. System Dynamics (State Equations)

4.1 Pennes Bioheat Equation (Continuous PDE)

The temperature field is governed by the Pennes bioheat equation:

ρ(r) c(r) ∂T/∂t = ∇·(k(r) ∇T) − ω_b(r) ρ_b c_b (T − T_b) + Q_met(r) + Q_source(r, u, t)

where the terms are defined in the following table:

- ρ(r) [kg/m³]: tissue mass density

- c(r) [J/(kg·K)]: tissue specific heat capacity

- k(r) [W/(m·K)]: thermal conductivity (possibly anisotropic tensor)

- ω_b(r) [s⁻¹]: blood perfusion rate

- ρ_b = 1050 kg/m³: blood density

- c_b = 3600 J/(kg·K): blood specific heat

- T_b [°C]: arterial blood temperature (body core, ~37°C)

- Q_met(r) [W/m³]: metabolic heat generation rate

- Q_source(r, u, t) [W/m³]: volumetric heat deposition from ablation source

4.2 Microwave Heat Source Model

For microwave ablation, the heat source is proportional to the Specific Absorption Rate (SAR):

Q_source(r, u, t) = SAR(r, q(t)) · P(t) = σ(r)|E(r, q(t))|² · P(t)

where σ(r) [S/m] is tissue electrical conductivity and E(r, q(t)) [V/m] is the normalized electric field pattern (precomputed from a finite element EM model). The spatial pattern b(r, q(t)) encodes the antenna near-field distribution:

Q_source(r, u, t) = b(r, q) · P(t)

4.3 Damage Dynamics

∂Ω_d/∂t = A · exp(−E_a / (R · T(r, t))),    r ∈ Ω, t ∈ (0, t_f]

4.4 Discretized State-Space Form

Applying finite difference or finite element spatial discretization (mesh size h, N voxels) yields the ODE system:

ẋ(t) = f(x(t), u(t)) = A_d x_T(t) + B(u(t)) + d(x_T(t))

where:

- A_d ∈ ℝᴺˣᴺ: sparse diffusion-perfusion matrix (tri- or penta-diagonal for FD, or FEM stiffness/mass matrix)

- B(u) ∈ ℝᴺ: control input vector, B_i = b(rᵢ, q)·P

- d(x_T) ∈ ℝᴺ: Arrhenius damage rate vector, dᵢ = A·exp(−E_a/(R·Tᵢ))

Explicitly:

ẋ_T = (1/(ρc)) [K_d x_T − W_b (x_T − T_b 1) + Q_met + B_P P(t)]

ẋ_Ω = A · exp(−E_a / (R · x_T))   [elementwise]

where K_d ∈ ℝᴺˣᴺ is the (FD or FEM) discrete Laplacian weighted by k(r), and W_b = diag(ω_b,i ρ_b c_b).

5. Boundary and Initial Conditions

5.1 Initial Conditions

T(r, 0) = T₀(r),    r ∈ Ω  [pre-treatment baseline temperature]

Ω_d(r, 0) = 0,       r ∈ Ω  [no pre-existing thermal damage]

Typically T₀(r) ≈ 37°C (body temperature) throughout the domain, or can be the measured pre-treatment MRI thermometry map.

5.2 Boundary Conditions on ∂Ω

Three standard options (select based on anatomy):

Dirichlet (Fixed Temperature — large vessel or body surface):

T(r, t) = T_∞,    r ∈ ∂Ω_D

Neumann (Insulated — symmetry plane):

k(r) ∂T/∂n = 0,    r ∈ ∂Ω_N

Robin (Convective Cooling — tissue-air or tissue-fluid interface):

k(r) ∂T/∂n + h_c(T − T_∞) = 0,    r ∈ ∂Ω_R

where h_c [W/(m²·K)] is the convective heat transfer coefficient and T_∞ is the ambient temperature.

6. Domain Partition

The tissue domain Ω is partitioned into three disjoint non-overlapping regions:

Ω = Ω_T ∪ Ω_H ∪ Ω_M,    Ω_T ∩ Ω_H = Ω_T ∩ Ω_M = Ω_H ∩ Ω_M = ∅

- Ω_T ⊂ Ω: tumor (target) region — to be fully ablated

- Ω_H ⊂ Ω: healthy tissue — to be protected

- Ω_M ⊂ Ω: safety margin / transition zone (optional)

These regions are defined from pre-treatment MRI segmentation.

7. Optimal Control Problem Formulation

7.1 Standard Bolza Form

The complete optimal control problem is stated in Bolza form:

min_{u(·), t_f}  J[u, t_f] = Φ(x(t_f), t_f) + ∫₀^t_f L(x(t), u(t), t) dt

subject to the state dynamics, boundary conditions, control constraints, and terminal constraints detailed below.

7.2 Running Cost L (Lagrangian)

The running cost captures three competing objectives during the treatment:

L(x, u, t) = α₁ ‖u(t)‖² + α₂ ∫_{Ω_H} max(0, T(r,t) − T_safe)² dr + α₃

- α₁ ‖u‖²: penalize excessive energy deposition (reduce tissue damage, device wear)

- α₂ ∫_{Ω_H} (T − T_safe)₊² dr: quadratic penalty for healthy tissue temperature exceeding threshold T_safe ≈ 42–45°C

- α₃ > 0: constant rate cost on time; encourages minimum-time treatment

7.3 Terminal Cost Φ (Mayer Term)

The Mayer term penalizes incomplete tumor ablation at the final time t_f:

Φ(x(t_f), t_f) = γ₁ ∫_{Ω_T} max(0, 1 − Ω_d(r, t_f))² dr + γ₂ t_f

- γ₁ ∫_{Ω_T} (1 − Ω_d)₊²: soft penalty for unablated tumor voxels

- γ₂ t_f: direct penalization of total procedure time

7.4 Complete Cost Functional

Combining the Lagrangian and Mayer terms:

J[u, t_f] = γ₁ ∫_{Ω_T} max(0, 1 − Ω_d(r, t_f))² dr + γ₂ t_f

+ ∫₀^t_f [ α₁‖u‖² + α₂ ∫_{Ω_H} max(0, T − T_safe)² dr + α₃ ] dt

The weights α₁, α₂, α₃, γ₁, γ₂ > 0 are tuning parameters reflecting clinical priorities.

8. Constraints

8.1 State Dynamics (Equality Constraint)

ẋ(t) = f(x(t), u(t)),    t ∈ (0, t_f]

x(0) = x₀ = [T₀ᵀ, 0ᵀ]ᵀ

8.2 Path (State-Control) Inequality Constraints

These must hold for all times t ∈ [0, t_f] and all spatial locations:

(a) Healthy tissue temperature limit:

T(r, t) ≤ T_safe,    ∀ r ∈ Ω_H, ∀ t ∈ [0, t_f]

(b) Absolute safety temperature (critical structures):

T(r, t) ≤ T_crit,    ∀ r ∈ Ω_crit ⊂ Ω_H, ∀ t ∈ [0, t_f]

(c) Control box constraints:

0 ≤ P(t) ≤ P_max,    ∀ t ∈ [0, t_f]

(d) Control rate constraints (actuator slew limit):

|dP/dt| ≤ Ṗ_max,    ∀ t ∈ [0, t_f]

8.3 Terminal (Endpoint) Constraints

(a) Tumor ablation requirement (hard constraint or soft via Φ):

Ω_d(r, t_f) ≥ 1,    ∀ r ∈ Ω_T

(b) Free final time positivity:

t_f > 0

(c) Maximum allowable treatment time:

t_f ≤ T_max

8.4 Summary: Complete Constrained Problem

min_{u(·) ∈ 𝒰, t_f ∈ (0, T_max]}  J[u, t_f]

subject to:

ẋ = f(x, u),    x(0) = x₀                   [dynamics]

T(r, t) ≤ T_safe,  ∀ r ∈ Ω_H, t ∈ [0, t_f]  [path safety]

0 ≤ P(t) ≤ P_max                             [control bounds]

Ω_d(r, t_f) ≥ 1,  ∀ r ∈ Ω_T               [terminal ablation]

9. Pontryagin Minimum Principle and Optimality Conditions

9.1 Hamiltonian

Define the costate (adjoint variable) λ(t) ∈ ℝ²ᴺ, partitioned as λ = [λ_T, λ_Ω]ᵀ. The Hamiltonian is:

ℋ(x, u, λ, t) = L(x, u, t) + λᵀ f(x, u)

= α₁‖u‖² + α₂‖max(0, T − T_safe)‖²_{Ω_H} + α₃ + λ_T^ᵀ ẋ_T + λ_Ω^ᵀ ẋ_Ω

9.2 Adjoint (Costate) Equations

The adjoint PDEs evolve backward in time from t = t_f to t = 0:

−λ̇_T = ∂ℋ/∂T = A_dᵀ λ_T − W_b λ_T + 2α₂ max(0, T − T_safe) 1_{Ω_H}

+ λ_Ω ⊙ [A (E_a)/(R T²) exp(−E_a/(R T))]

−λ̇_Ω = ∂ℋ/∂Ω_d = 0    ⟹   λ_Ω(t) = λ_Ω(t_f) = const

9.3 Transversality Conditions (Free Final Time)

Since t_f is free, the Hamiltonian must vanish at the terminal time:

ℋ(x(t_f), u(t_f), λ(t_f), t_f) + ∂Φ/∂t_f = 0

⟹  L(x(t_f), u(t_f), t_f) + λ(t_f)ᵀ f(x(t_f), u(t_f)) + γ₂ = 0

The terminal costate conditions (complementary slackness for inequality constraints) are:

λ_T(t_f) = ∂Φ/∂x_T|_{t_f} = 0  [no terminal cost on T]

λ_Ω(t_f) = ∂Φ/∂x_Ω|_{t_f} = −2γ₁ max(0, 1 − Ω_d(r, t_f)) 1_{Ω_T}

9.4 Optimal Control Law (Pontryagin)

The PMP states that the optimal control minimizes the Hamiltonian pointwise:

u*(t) = argmin_{u ∈ 𝒰}  ℋ(x*(t), u, λ*(t), t)

For the quadratic power term α₁ P² and linear control-to-state coupling:

∂ℋ/∂P = 2α₁ P − λ_Tᵀ B_P = 0  ⟹  P*(t) = (1/2α₁) λ_Tᵀ B_P,  projected onto [0, P_max]

P*(t) = max(0, min(P_max, λ_Tᵀ B_P / (2α₁)))

10. Problem Classification

- PDE-constrained optimal control (distributed parameter system)

- Infinite-dimensional state space (before discretization): x ∈ L²(Ω) × L²(Ω)

- Finite-dimensional after discretization: x ∈ ℝ²ᴺ, u ∈ ℝᵐ

- Free final time: t_f is an optimization variable

- Bolza problem: running cost L + terminal cost Φ

- State-constrained: path constraints on T(r,t) in Ω_H

- Control-constrained: 0 ≤ P ≤ P_max, |Ṗ| ≤ Ṗ_max

- Soft terminal equality (via Φ) or hard inequality: Ω_d(t_f) ≥ 1 in Ω_T

- Nonlinear dynamics: Arrhenius damage is exponentially nonlinear in T

- Possibly hybrid: tissue properties may switch after necrosis (discontinuous dynamics)

11. Discretized Finite-Dimensional Formulation (Direct Method)

For numerical solution, the PDE-constrained problem is converted to a finite-dimensional NLP via direct transcription (e.g., direct collocation or direct multiple shooting).

11.1 Spatial Discretization (Finite Difference)

On a uniform Cartesian grid with spacing h, the discrete diffusion-perfusion matrix is:

A_d = (k/h²) L_FD − W_b,    L_FD = standard discrete Laplacian ∈ ℝᴺˣᴺ

The discrete system becomes:

M ẋ_T = A_d x_T + b_P P(t) + Q_met

where M = diag(ρᵢ cᵢ) is the mass matrix and b_P ∈ ℝᴺ is the normalized SAR distribution vector.

11.2 Temporal Discretization (Collocation)

Dividing [0, t_f] into K intervals with nodes t_k = k·Δt, and applying e.g. trapezoidal collocation:

x_{k+1} = x_k + (Δt/2)[f(x_k, u_k) + f(x_{k+1}, u_{k+1})]

This converts the infinite-dimensional problem into a large-scale NLP in the decision variables {x_k, u_k}_{k=0}^{K} and t_f.

11.3 NLP Form

min_{x_{0:K}, u_{0:K}, t_f}  ∑_k L(x_k, u_k)Δt + Φ(x_K, t_f)

s.t.  x_{k+1} = x_k + (Δt/2)[f(x_k, u_k) + f(x_{k+1}, u_{k+1})]   ∀k

x_0 = x_init

g(x_k, u_k) ≤ 0   ∀k  [path constraints]

h(x_K) ≤ 0        [terminal constraints]

u_k ∈ 𝒰            ∀k

12. Parameter Summary

Table 1 below summarizes key physical parameters and their typical values for soft tissue ablation:

- ρ: tissue density — 1000–1100 kg/m³

- c: specific heat — 3400–3800 J/(kg·K)

- k: thermal conductivity — 0.4–0.6 W/(m·K)

- ω_b: blood perfusion rate — 0.003–0.01 s⁻¹

- T_b: blood temperature — 37°C

- A (Arrhenius): frequency factor — 3.1×10⁹⁸ s⁻¹ (for protein denaturation)

- E_a (Arrhenius): activation energy — 6.28×10⁵ J/mol

- R: universal gas constant — 8.314 J/(mol·K)

- T_safe: healthy tissue safety threshold — 42–45°C

- T_crit: critical structure absolute limit — 40°C

- P_max: maximum applicator power — 30–100 W

- T_max: maximum allowed procedure time — 10–30 min

13. Solution Approaches

13.1 Indirect Methods (PMP-Based)

- Solve the two-point boundary value problem (TPBVP) formed by state + adjoint PDEs

- Forward integration of state, backward integration of adjoint

- Shooting methods: single shooting, multiple shooting

- Convergence challenges due to nonlinearity and state constraints

13.2 Direct Methods (Transcription to NLP)

- Direct collocation (Hermite-Simpson, Radau, Gauss-Legendre)

- Direct multiple shooting

- Sparse NLP solvers: IPOPT, SNOPT, KNITRO

- Scales to N ≈ 10³–10⁴ voxels with structure exploitation

13.3 Model Predictive Control (Receding Horizon)

- Solve the OCP repeatedly over a rolling prediction horizon T_p

- Apply the first control action, then re-solve with updated MRI measurement

- Naturally handles measurement feedback and model uncertainty

- Critical for real-time MRI-guided closed-loop control

13.4 Reinforcement Learning / Data-Driven

- Policy gradient or actor-critic methods on the discretized MDP

- Can learn from simulation (PDE solver) without analytical gradients

- Suited for handling imaging noise and patient variability

14. Extended Formulation: Robotic MRI-Guided Ablation

For a robotic system that actuates probe position and orientation inside the MRI bore, the state and dynamics are augmented:

14.1 Augmented State

x_aug(t) = [x_robot(t)ᵀ, x_T(t)ᵀ, x_Ω(t)ᵀ]ᵀ ∈ ℝ^(n_q + 2N)

where x_robot = [q(t)ᵀ, q̇(t)ᵀ]ᵀ ∈ ℝ^(2n_q) is the robot kinematic state.

14.2 Coupled Dynamics

q̈ = M_r(q)⁻¹ [τ(t) − C(q, q̇)q̇ − G(q)]         [robot dynamics]

ẋ_T = f_T(x_T, P(t), q(t))                        [bioheat with probe position]

ẋ_Ω = f_Ω(x_T)                                    [Arrhenius damage]

14.3 Augmented Control

u_aug(t) = [τ(t)ᵀ, P(t)]ᵀ ∈ ℝ^(n_q+1)

where τ(t) ∈ ℝⁿq are the robot joint torques (or equivalent actuator commands).

The probe position p(q) ∈ ℝ³ enters the SAR field via the forward kinematics, making the bioheat source term Q_source explicitly dependent on q(t), creating a fully coupled, nonlinear PDE-ODE optimal control problem.

References and Further Reading

- Pennes, H.H. (1948). Analysis of tissue and arterial blood temperatures in the resting human forearm. J. Appl. Physiol., 1(2), 93–122.

- Moritz, A.R., Henriques, F.C. (1947). Studies of thermal injury II. Am. J. Pathol., 23(5), 695–720. [Arrhenius damage model]

- Bett, S., et al. (2022). Model predictive control for MRI-guided thermal ablation. IEEE Trans. Biomed. Eng.

- Liberzon, D. (2012). Calculus of Variations and Optimal Control Theory. Princeton University Press.

- Biegler, L.T. (2010). Nonlinear Programming: Concepts, Algorithms, and Applications to Chemical Processes. SIAM.

- Kirk, D.E. (1970). Optimal Control Theory: An Introduction. Dover.